| URGENT FIELD SAFETY NOTICE |  |
| --- | --- |

| Date of Letter Deployment | GE HealthCare Ref. # 40910 |
| --- | --- |

| To: |  | Director of Clinical/Radiology Risk Manager/Hospital Administrator Director of Biomedical Engineering |  |  |
| --- | --- | --- | --- | --- |
| RE: |  | Certain Infinia, VariCam/VG/VH, Brivo NM 615, Discovery NM 630, Discovery NM/CT 670, and Optima NM/CT 640 Nuclear Medicine systems |  |  |
| Safety Issue |  | GE HealthCare has become aware that certain Infinia, VariCam/VG/VH, Brivo NM 615, Discovery NM 630, Discovery NM/CT 670, and Optima NM/CT 640 Nuclear Medicine systems (see Affected Products Details and Appendices) could have been transported or relocated without proper detector support, which can result in excessive load being applied to the detector mounting mechanisms. If this occurs, it can potentially result in a detector fall and lead to life-threatening bodily injury. - There have been no reports of detector falls or injuries as a result of this issue. |

| Actions to be taken by Customer/User | - There are two populations potentially affected by this notification: - For Serial Numbers listed in Appendix A of this letter: - GE HealthCare requests that you STOP USING the system until GE HealthCare has inspected and if necessary, corrected your system and provided you with a letter indicating the system can be returned to clinical use. GE HealthCare has determined that these systems could have previously been transported or relocated without proper detector support. - For Serial Numbers listed in Appendix B of this letter: - You can continue to use your system, as GE HealthCare has determined that these systems have not previously been transported or relocated without proper detector support. - For Serial Numbers listed in either Appendix A or Appendix B of this letter: - GE HealthCare requests that you DO NOT physically relocate or transport your system from its current location. If you need to move or transport your system, please email GE HealthCare at recall.40910@gehealthcare.com to receive specific instructions to follow before moving the system. When emailing this address, please provide the following information: - System serial number - Current system location - The reason for the move or transport, and subsequent destination - Do not proceed with a physical relocation or transport of your system until you have received these instructions and taken the actions described to avoid compromising the detector mounting mechanism. - Please ensure all potential users in your facility are made aware of this safety notification and the requested actions. - Please complete and return the attached acknowledgement form to recall.40910@gehealthcare.com - Please retain this document for your records. |
| --- | --- |
| Affected Product Details | The following Nuclear Medicine systems are potentially affected: - System Type; - Affected Systems. - Infinia Nuclear Medicine Systems: Infinia Infinia Hawkeye Infinia II Infinia II Hawkeye 4; - See Appendices A and B for serial numbers. - VG Nuclear Medicine systems: Varicam Millennium VG Millennium VG Hawkeye Discovery VH; - See Appendices A and B for serial numbers. - Brivo NM 615; - See Appendices A and B for serial numbers. - Discovery NM 630; - See Appendices A and B for serial numbers. - Discovery NM/CT 670 Discovery NM/CT 670 ES Discovery NM/CT 670 DR Discovery NM/CT 670 PRO Discovery NM/CT 670 CZT; - See Appendices A and B for serial numbers. - Optima NM/CT 640; - See Appendices A and B for serial numbers. - To identify the serial number of your system, please refer to the “rating plate” located according to the illustration below (see Figure 1 for location of the “rating plate” for Infinia systems and Figure 2 for location of the “rating plate” for NM600 systems): - : - Figure 1: Picture of Infinia with location of rating plate - Figure 2: Picture of NM600 with location of rating plate - Intended Use: Infinia, VariCam: The system is intended to be used to perform general Nuclear Medicine imaging procedures for detection of radioisotope tracer uptake in the patient body, using a variety of scanning modes supported by various acquisition types and optional imaging features designed to enhance image quality in Oncology, Cardiology, Neurology and other clinical diagnostic imaging applications. Brivo NM 615: The Brivo NM 615 system is intended for General Nuclear Medicine imaging procedures using variety of scanning modes supported by various acquisition types. This generic type of device can include signal analysis and display equipment, patient and equipment support, and accessories. Optima NM/CT 640: The GE Optima NM/CT 640 system is intended for General Nuclear Medicine imaging procedures for detection of radioisotope tracer uptake in the patient body. It includes a general purpose Nuclear Medicine (NM) system using a variety of scanning modes supported by various acquisition types, and a CT component which is intended specifically for enabling attenuation correction and anatomical localization on SPECT studies. - Discovery NM 630: The GE Discovery NM 630 system is a medical tool intended for use by appropriately trained healthcare professionals to aid in detecting, localizing, diagnosing of diseases and in assessment of organ function for the evaluation of diseases, trauma, abnormalities, and disorders such as, but not limited to, cardiovascular disease, neurological disorders and cancer. The system output can also be used by the physician for staging and restaging of tumors, planning, guiding, and monitoring therapy Discovery NM/CT 670, 670 ES\DR\PRO\CZT: The GE HealthCare Discovery NM/CT 670 system is intended for general Nuclear Medicine imaging procedures for detection of radioisotope tracer uptake in the patient body. It includes a general purpose Nuclear Medicine system using a variety of scanning modes supported by various acquisition types, and a Computed Tomography system which is intended for enabling attenuation correction and anatomical localization of SPECT images and for standalone head, whole body, cardiac and vascular X-ray Computed Tomography applications. |
| Product Correction | For serial numbers listed in Appendix A of this letter: A GE HealthCare representative will contact you to schedule an appointment as soon as possible. During the scheduled appointment, a GE HealthCare Field Engineer will inspect the potentially affected system to determine if the detector support mechanisms are intact and functioning correctly. - If an issue is found with the detector mounting mechanism during inspection, your system will be taken out of service and an appointment will be made to correct the system, at no cost to you. We will work diligently to conduct this activity as quickly as possible to minimize any disruption to clinical care. - If the detector mounting mechanism is found not to be affected, the GE HealthCare Field Engineer will provide you with a letter, indicating the system can be returned to clinical use. For serial numbers listed in either Appendix A or B of this letter: GE HealthCare will apply warning labels to your device and will provide you with updated instructions on how to correctly relocate your system, so the detector mounting mechanism is not compromised. |
| Contact Information | If you have any questions or concerns regarding this notification, please contact GE HealthCare Service or your local Service Representative. |

GE Healthcare confirms that this notice has been notified to the appropriate Regulatory Agency.

Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us per the contact information above.

Sincerely,

Laila Gurney 						Scott Kelley

Chief Quality & Regulatory Officer 			Chief Medical Officer

GE Healthcare 						GE Healthcare

|  |  | GE HealthCare Ref. # 40910 |
| --- | --- | --- |

MEDICAL DEVICE NOTIFICATION ACKNOWLEDGEMENT

RESPONSE REQUIRED

Please complete this form and return it to GE HealthCare promptly upon receipt and no later than 30 days from receipt. This will confirm receipt and understanding of the Medical Device Correction Notice.

| Facility Name: |  |
| --- | --- |
| Street Address: |  |
| City/State/ZIP/Country: |  |
| Customer Email Address: |  |
| Customer Phone Number: |  |

| ☐ | We acknowledge that we have stopped use of the affected Nuclear Medicine system(s) and will not resume use until an inspection is completed by GE HealthCare and GE HealthCare has indicated the system(s) can return to clinical use. We acknowledge receipt and understanding of the accompanying Urgent Medical Device Correction Letter, and that we have informed appropriate staff and have taken and will take appropriate actions in accordance with that Notification. |
| --- | --- |

Please provide the name of the individual with responsibility who completed this form.

| Signature: |  |
| --- | --- |
| Printed Name: |  |
| Position/Job Title: |  |
| Date (DD/MM/YYYY): |  |

Please return completed form by scanning or taking a photo of the completed form and email to recall.40910@gehealthcare.com

Appendix A – Impacted Serial Numbers, STOP USING your system: Infinia and Infinia Hawkeye Model Serial Numbers. 16002; 16328; 16614; 16839; 17108; 17296; 17525; 17726. 16003; 16330; 16621; 16841; 17122; 17299; 17536; 17727. 16008; 16334; 16622; 16845; 17129; 17301; 17538; 17728. 16010; 16342; 16634; 16851; 17132; 17304; 17540; 17729. 16011; 16349; 16638; 16854; 17138; 17306; 17543; 17730. 16012; 16356; 16640; 16874; 17142; 17307; 17551; 17731. 16013; 16372; 16642; 16878; 17145; 17313; 17556; 17732. 16016; 16381; 16646; 16884; 17154; 17324; 17557; 17733. 16024; 16388; 16647; 16889; 17158; 17327; 17569; 17735. 16032; 16406; 16649; 16894; 17159; 17328; 17574; 17736. 16065; 16411; 16661; 16896; 17161; 17329; 17587; 17737. 16074; 16419; 16667; 16899; 17168; 17348; 17588; 17741. 16079; 16429; 16672; 16904; 17172; 17355; 17591; 17745. 16081; 16431; 16679; 16906; 17181; 17359; 17610; 17749. 16087; 16445; 16681; 16908; 17183; 17370; 17611; 17751. 16106; 16447; 16684; 16915; 17186; 17372; 17614; 17752. 16117; 16465; 16690; 16937; 17192; 17375; 17617; 17757. 16126; 16467; 16696; 16938; 17202; 17377; 17625; 17760. 16138; 16469; 16704; 16939; 17212; 17382; 17630; 17765. 16142; 16474; 16710; 16941; 17213; 17393; 17633; 17775. 16150; 16475; 16719; 16960; 17214; 17395; 17639; 17789. 16167; 16477; 16732; 16965; 17216; 17402; 17640; 17790. 16173; 16485; 16733; 16992; 17217; 17407; 17644; 17791. 16186; 16504; 16761; 16996; 17221; 17411; 17645; 17793. 16190; 16513; 16767; 16997; 17226; 17413; 17649; 17794. 16196; 16531; 16772; 17008; 17238; 17419; 17657; 17795. 16207; 16535; 16774; 17009; 17239; 17421; 17658; 17801. 16214; 16538; 16784; 17016; 17243; 17427; 17659; 17808. 16221; 16539; 16785; 17020; 17257; 17428; 17671; 17810. 16229; 16542; 16786; 17032; 17260; 17432; 17672; 17813. 16235; 16551; 16788; 17036; 17262; 17467; 17673; 17814. 16259; 16554; 16792; 17042; 17265; 17476; 17685; 17816. 16270; 16562; 16802; 17043; 17266; 17493; 17686; 17818. 16278; 16566; 16803; 17050; 17270; 17494; 17689; 17819. 16279; 16573; 16811; 17064; 17273; 17507; 17690; 17821. 16319; 16577; 16814; 17066; 17279; 17509; 17691; 17822. 16322; 16587; 16815; 17073; 17281; 17518; 17694; 17825. 16325; 16600; 16820; 17081; 17284; 17521; 17712; 17832. 16326; 16606; 16827; 17092; 17289; 17522; 17713; 17840. 16327; 16613; 16837; 17107; 17290; 17523; 17719; 17841. Appendix A – Impacted Serial Numbers, STOP USING your system: Infinia and Infinia Hawkeye Model Serial Numbers. 17843; 17970; 18121; 18279; 18480; 18570; 18615; 18151. 17849; 17972; 18124; 18280; 18484; 18571; 18616; 18179. 17850; 17977; 18125; 18283; 18490; 18572; 18617; 18182. 17853; 17980; 18126; 18286; 18492; 18573; 18618; 18247. 17854; 17988; 18129; 18289; 18497; 18574; 18619; 18393. 17855; 17993; 18130; 18292; 18502; 18575; 18620; 18431. 17856; 17996; 18132; 18295; 18510; 18577; 18622; 18468. 17857; 18001; 18135; 18299; 18512; 18578; 18624; 18488. 17859; 18002; 18136; 18301; 18517; 18579; 18627; 18493. 17860; 18007; 18142; 18306; 18519; 18580; 18628; 18495. 17861; 18009; 18148; 18310; 18522; 18581; 18629; 18504. 17863; 18010; 18149; 18318; 18523; 18582; 18630; 18635. 17864; 18015; 18156; 18322; 18525; 18583; 18631. 17866; 18027; 18170; 18332; 18529; 18584; 18632. 17871; 18042; 18173; 18339; 18533; 18585; 18633. 17873; 18047; 18187; 18352; 18537; 18586; 18634. 17874; 18049; 18192; 18362; 18538; 18587; 18637. 17882; 18056; 18194; 18363; 18542; 18588; 18638. 17884; 18059; 18203; 18364; 18543; 18589; 18639. 17888; 18060; 18206; 18366; 18544; 18590; 18640. 17890; 18063; 18209; 18373; 18545; 18591; 18641. 17897; 18067; 18213; 18375; 18546; 18592; 18643. 17901; 18073; 18223; 18377; 18547; 18593; 18644. 17907; 18076; 18226; 18387; 18551; 18595; 18645. 17908; 18080; 18227; 18389; 18552; 18596; 18646. 17910; 18081; 18233; 18397; 18553; 18597; 18647. 17912; 18084; 18234; 18410; 18554; 18599; 18648. 17916; 18086; 18237; 18413; 18556; 18600; 18649. 17918; 18087; 18240; 18414; 18557; 18601; 18650. 17919; 18089; 18242; 18422; 18558; 18602; 18651. 17920; 18093; 18243; 18425; 18559; 18604; 18652. 17922; 18095; 18248; 18426; 18560; 18605; 18653. 17926; 18097; 18252; 18430; 18561; 18606; 18654. 17937; 18099; 18253; 18435; 18562; 18607; 18655. 17958; 18100; 18257; 18436; 18563; 18608; 18656. 17962; 18104; 18259; 18440; 18564; 18609; 18657. 17964; 18109; 18266; 18442; 18565; 18610; 19191. 17965; 18114; 18271; 18455; 18567; 18611; 17126. 17967; 18115; 18275; 18458; 18568; 18613; 17768. 17969; 18119; 18278; 18478; 18569; 18614; 18137. Varicam, VH and VG Model Serial Numbers. 10550; 10569; 10570; 10582. 10587; 10591; 10595; 1092. Appendix A – Impacted Serial Numbers, STOP USING your system: 600 Series Rotor A Model Serial Numbers. 21001; 21053; 21103; 21151; 21199; 21253; 21312; 21362. 21002; 21056; 21104; 21152; 21200; 21255; 21313; 21363. 21003; 21058; 21105; 21154; 21201; 21258; 21314; 21364. 21004; 21059; 21106; 21155; 21202; 21259; 21315; 21365. 21005; 21060; 21107; 21156; 21203; 21260; 21316; 21366. 21006; 21061; 21108; 21157; 21204; 21262; 21317; 21367. 21007; 21062; 21109; 21158; 21205; 21263; 21318; 21368. 21009; 21063; 21110; 21159; 21207; 21265; 21319; 21369. 21010; 21064; 21111; 21160; 21209; 21266; 21320; 21370. 21011; 21065; 21113; 21161; 21211; 21267; 21322; 21371. 21012; 21066; 21114; 21162; 21212; 21269; 21323; 21375. 21013; 21067; 21115; 21163; 21213; 21270; 21324; 21376. 21016; 21068; 21118; 21164; 21214; 21271; 21325; 21377. 21017; 21070; 21119; 21165; 21215; 21272; 21326; 21378. 21018; 21071; 21120; 21166; 21217; 21273; 21327; 21379. 21019; 21072; 21121; 21167; 21218; 21274; 21328; 21380. 21020; 21073; 21122; 21168; 21219; 21275; 21329; 21381. 21021; 21074; 21123; 21169; 21220; 21276; 21330; 21382. 21023; 21075; 21124; 21170; 21221; 21279; 21331; 21383. 21024; 21076; 21125; 21171; 21222; 21280; 21332; 21384. 21025; 21077; 21126; 21172; 21223; 21281; 21333; 21385. 21026; 21078; 21127; 21173; 21224; 21283; 21334; 21386. 21027; 21079; 21128; 21174; 21225; 21284; 21335; 21387. 21028; 21080; 21129; 21175; 21226; 21285; 21336; 21388. 21029; 21081; 21130; 21176; 21229; 21286; 21337; 21389. 21030; 21082; 21131; 21177; 21230; 21287; 21338; 21390. 21031; 21083; 21132; 21178; 21231; 21288; 21339; 21391. 21032; 21084; 21133; 21180; 21233; 21289; 21340; 21392. 21033; 21085; 21134; 21182; 21234; 21290; 21344; 21393. 21035; 21087; 21135; 21183; 21235; 21291; 21345; 21395. 21037; 21088; 21136; 21184; 21236; 21292; 21346; 21396. 21039; 21089; 21137; 21185; 21237; 21294; 21347; 21397. 21040; 21090; 21138; 21186; 21239; 21295; 21348; 21398. 21041; 21091; 21139; 21187; 21240; 21296; 21349; 21399. 21042; 21092; 21140; 21188; 21241; 21298; 21351; 21400. 21043; 21093; 21141; 21189; 21242; 21299; 21352; 21401. 21044; 21094; 21142; 21190; 21243; 21300; 21353; 21402. 21045; 21095; 21143; 21191; 21244; 21301; 21354; 21403. 21046; 21096; 21144; 21192; 21245; 21302; 21355; 21404. 21047; 21097; 21145; 21193; 21246; 21304; 21356; 21405. 21048; 21098; 21146; 21194; 21247; 21305; 21357; 21406. 21049; 21099; 21147; 21195; 21248; 21307; 21358; 21407. 21050; 21100; 21148; 21196; 21250; 21309; 21359; 21408. 21051; 21101; 21149; 21197; 21251; 21310; 21360; 21409. 21052; 21102; 21150; 21198; 21252; 21311; 21361; 21410. Appendix A – Impacted Serial Numbers, STOP USING your system: 600 Series Rotor A Model Serial Numbers. 21411; 22049; 22095; 22157; 22222; 22283; 22344; 22408. 21412; 22050; 22096; 22158; 22223; 22284; 22345; 22411. 21413; 22051; 22097; 22160; 22224; 22285; 22347; 22412. 22001; 22052; 22098; 22161; 22225; 22287; 22348; 22413. 22002; 22053; 22099; 22162; 22228; 22288; 22349; 22415. 22004; 22054; 22100; 22164; 22229; 22289; 22352; 22416. 22005; 22055; 22101; 22165; 22230; 22290; 22353; 22419. 22007; 22056; 22102; 22167; 22231; 22291; 22354; 22420. 22008; 22057; 22103; 22168; 22232; 22293; 22355; 22421. 22009; 22058; 22105; 22169; 22233; 22294; 22356; 22423. 22010; 22059; 22110; 22170; 22234; 22295; 22357; 22426. 22011; 22060; 22111; 22171; 22236; 22296; 22358; 22427. 22012; 22061; 22112; 22172; 22237; 22298; 22359; 22430. 22013; 22062; 22114; 22173; 22238; 22299; 22362; 22431. 22014; 22063; 22115; 22176; 22239; 22301; 22364; 22433. 22015; 22064; 22116; 22177; 22240; 22303; 22366; 22436. 22016; 22065; 22117; 22179; 22241; 22304; 22367; 22437. 22017; 22066; 22118; 22180; 22242; 22305; 22368; 22438. 22018; 22067; 22119; 22181; 22243; 22306; 22369; 22439. 22020; 22068; 22120; 22182; 22244; 22308; 22370; 22442. 22022; 22069; 22121; 22183; 22245; 22309; 22371; 22443. 22023; 22070; 22122; 22184; 22246; 22310; 22375; 22444. 22024; 22071; 22123; 22185; 22247; 22311; 22376; 22447. 22025; 22072; 22124; 22187; 22248; 22313; 22377; 22448. 22026; 22073; 22125; 22189; 22249; 22314; 22378; 22449. 22027; 22074; 22127; 22191; 22252; 22315; 22379; 22451. 22028; 22075; 22128; 22194; 22254; 22316; 22380; 22452. 22029; 22076; 22129; 22198; 22256; 22317; 22383; 22453. 22030; 22077; 22130; 22199; 22257; 22319; 22385; 22454. 22032; 22078; 22131; 22201; 22259; 22320; 22386; 22456. 22034; 22079; 22132; 22202; 22260; 22322; 22387; 22457. 22035; 22080; 22133; 22205; 22261; 22323; 22388; 22458. 22036; 22081; 22134; 22206; 22262; 22325; 22389; 22463. 22037; 22082; 22135; 22207; 22265; 22326; 22391; 22464. 22038; 22083; 22136; 22208; 22266; 22327; 22392; 22465. 22039; 22085; 22138; 22209; 22267; 22329; 22393; 22466. 22040; 22086; 22141; 22210; 22268; 22330; 22394; 22468. 22041; 22087; 22144; 22211; 22273; 22332; 22398; 22469. 22042; 22088; 22149; 22212; 22274; 22333; 22401; 22470. 22043; 22089; 22150; 22214; 22276; 22335; 22402; 22471. 22044; 22090; 22151; 22216; 22277; 22336; 22403; 22473. 22045; 22091; 22152; 22218; 22278; 22338; 22404; 22474. 22046; 22092; 22153; 22219; 22279; 22340; 22405; 22475. 22047; 22093; 22155; 22220; 22280; 22342; 22406; 22476. 22048; 22094; 22156; 22221; 22281; 22343; 22407; 22477. Appendix A – Impacted Serial Numbers, STOP USING your system: 600 Series Rotor A Model Serial Numbers. 22478; 22566; 24025; 25020; 26068; 26138; 26208; 26290. 22479; 22568; 24027; 26002; 26069; 26141; 26210; 26292. 22480; 22571; 24028; 26003; 26071; 26142; 26212; 26295. 22481; 22574; 24029; 26004; 26072; 26143; 26213; 26297. 22482; 22575; 24030; 26005; 26073; 26144; 26214; 26298. 22483; 22576; 24031; 26007; 26074; 26145; 26215; 26300. 22484; 22577; 24032; 26008; 26075; 26150; 26217; 26303. 22488; 22578; 24033; 26009; 26076; 26151; 26219; 26305. 22489; 22580; 24034; 26010; 26077; 26153; 26222; 26306. 22493; 22582; 24035; 26011; 26078; 26154; 26224; 26307. 22494; 22583; 24036; 26012; 26079; 26155; 26225; 26308. 22496; 22585; 24037; 26014; 26081; 26156; 26226; 26311. 22497; 22587; 24038; 26016; 26082; 26158; 26228; 26313. 22499; 22589; 24039; 26017; 26084; 26160; 26230; 26314. 22502; 22593; 24040; 26018; 26085; 26163; 26233; 26318. 22503; 22596; 24041; 26019; 26086; 26164; 26234; 26319. 22504; 22597; 24042; 26020; 26087; 26167; 26237; 26323. 22505; 22599; 24043; 26021; 26092; 26169; 26238; 26327. 22507; 22601; 24044; 26024; 26094; 26170; 26239; 26328. 22508; 22602; 24045; 26095; 26171; 26240; 26329. 22510; 22603; 24046; 26026; 26096; 26172; 26241; 26331. 22511; 22604; 24047; 26028; 26097; 26173; 26242; 26332. 22512; 22605; 24048; 26029; 26100; 26175; 26248; 26335. 22514; 22606; 24049; 26031; 26104; 26176; 26249; 26336. 22519; 22607; 24050; 26032; 26105; 26177; 26250; 26338. 22521; 24001; 24051; 26037; 26106; 26178; 26252; 26340. 22525; 24002; 24052; 26038; 26107; 26180; 26255; 26342. 22531; 24003; 24053; 26039; 26108; 26182; 26256; 26346. 22533; 24004; 24054; 26042; 26110; 26183; 26257; 26348. 22537; 24005; 24055; 26045; 26111; 26184; 26258; 26349. 22538; 24006; 24057; 26046; 26112; 26186; 26259; 26351. 22539; 24007; 24058; 26048; 26115; 26187; 26260; 27002. 22540; 24008; 24059; 26051; 26117; 26188; 26263; 27004. 22543; 24010; 24062; 26052; 26119; 26191; 26267; 27006. 22544; 24011; 24063; 26054; 26121; 26192; 26269; 27009. 22548; 24012; 24065; 26055; 26123; 26193; 26271; 27010. 22549; 24013; 24067; 26058; 26124; 26194; 26272; 27012. 22551; 24014; 24068; 26059; 26127; 26198; 26274; 27013. 22554; 24015; 24069; 26060; 26128; 26199; 26275; 27014. 22555; 24016; 24071; 26061; 26129; 26200; 26276; 27015. 22556; 24017; 24072; 26062; 26130; 26201; 26277; 27017. 22557; 24018; 24073; 26063; 26132; 26202; 26278; 27019. 22561; 24019; 24074; 26064; 26134; 26203; 26287; 27020. 22562; 24020; 24075; 26066; 26136; 26205; 26288; 27021. 22563; 24021; 24076; 26067; 26137; 26207; 26289; 27022. Appendix A – Impacted Serial Numbers, STOP USING your system: 600 Series Rotor A Model Serial Numbers. 27023; 27091; 630V22614; 630W22770; 630X22912; DDRW34022. 27028; 27093; 630V22618; 630W22771; 630X22913; DDRW34026. 27030; 27094; 630V22619; 630W22780; 630X22915; DDRW34027. 27032; 27095; 630V22620; 630W22781; 630X22920; DDRW34029. 27033; 27096; 630V22623; 630W22784; 670V21414; DDRW34030. 27034; 27097; 630V22624; 630W22791; 670V21415; DDRW34034. 27035; 27099; 630V22634; 630W22799; 670V21416; DDRW34035. 27036; 27101; 630V22635; 630W22803; 670V21417; DDRX34036. 27037; 27104; 630V22636; 630W22805; 670V21418; DDRX34037. 27039; 27105; 630V22643; 630W22810; 670V21419; DDRX34043. 27040; 27106; 630V22645; 630W22813; 670V21420; DDRX34051. 27043; 27107; 630V22646; 630W22826; 670V21421; DDRX34052. 27044; 27108; 630V22649; 630W22828; 670V21422; DDRX34054. 27045; 27109; 630V22650; 630W22829; 670V21423; DESV28014. 27046; 27111; 630V22657; 630W22834; 670V21424; DESV28016. 27047; 27112; 630V22661; 630W22836; 670V21425; DESV28017. 27050; 27113; 630V22667; 630W22841; 670W21427; DESV28018. 27052; 27114; 630V22673; 630W22844; 670W21428; DESW28023. 27053; 27116; 630V22678; 630W22847; 670W21429; DESW28026. 27054; 27120; 630V22679; 630W22853; 670W21430; DESW28027. 27055; 27122; 630V22680; 630W22854; 670W21431; DESW28028. 27057; 27125; 630V22682; 630W22855; 670W21432; DESW28030. 27059; 27130; 630V22683; 630W22861; 670W21433; DESX28031. 27061; 27132; 630V22684; 630W22866; 670W21434; NT24G1600020. 27062; 27134; 630V22685; 630W22867; 670W21435; NT24G1600021. 27063; 27135; 630V22687; 630X22868; 670W21436; NT24G1600022. 27066; 27136; 630V22692; 630X22869; 670W21437; NT24G1600023. 27067; 28002; 630V22693; 630X22877; 670W21438; NT24G1700001. 27068; 28005; 630V22697; 630X22881; 670W21439; NT24G1700002. 27069; 28006; 630V22698; 630X22882; 670W21440; NT24G1700003. 27070; 28007; 630V22704; 630X22884; 670W21441; NT24G1700004. 27073; 28008; 630V22705; 630X22885; 670W21442; NT24G1700005. 27074; 28009; 630V22711; 630X22886; 85748HL2; NT24G1700006. 27076; 28010; 630V22718; 630X22887; CZTV29006; NT24G1700007. 27078; 28011; 630W22720; 630X22888; CZTW29008; NT24G1700008. 27079; 28012; 630W22721; 630X22893; DDRW34001; NT24G1700009. 27081; 29003; 630W22722; 630X22895; DDRW34002; NT24G1700010. 27082; 153537HL6; 630W22723; 630X22897; DDRW34008; NT24G1700010BH. 27083; 615V24077; 630W22737; 630X22900; DDRW34009; NT24G1800001. 27084; 615W24078; 630W22742; 630X22901; DDRW34011; NT24G1800002. 27085; 615W24079; 630W22746; 630X22903; DDRW34012; NT24G1800003. 27086; 615W24080; 630W22747; 630X22904; DDRW34014; NT24G1800004. 27087; 630V22610; 630W22753; 630X22908; DDRW34017; NT24G1800005. 27088; 630V22611; 630W22764; 630X22909; DDRW34019; NT24G1800006. 27090; 630V22613; 630W22766; 630X22911; DDRW34021; NT24G1800007. Appendix A – Impacted Serial Numbers, STOP USING your system: 600 Series Rotor A Model Serial Numbers. NT26G1800001; ONMW26406; T26G1600030/144953HL7. NT26G1800002; ONMW26412; T26G1600031/136767HL1. NT26G1800003; ONMW26419; T26G1600032. NT26G1800004; ONMW26421; T26G1600033/145547HL6. NT26G1800005; ONMW26427; T26G1600034/136767HL1. NT26G1800008; ONMW26432; T26G1600035/147228HL1. NT26G1800009; ONMW26434; T26G16020/133975HL3. NT26G1800010; ONMW26438; T26G16021/133975HL3. NT26G1800011; ONMW26440; T26G16022/135526HL2. NT26G1800012; ONMW26445; T26G16023/135745HL8. NT26G1800013; ONMW26448; T26G16024/136251HL6. NT28G1800002; ONMW26449; T26G16025. NT28G1800003; ONMW26453; T26G1700036/148833HL7. NT28G1800004; ONMW26455; T26G1700037/152214HL3. NT2GG1800001; ONMW26456; T26G1700038/152744HL9. NT2JG1800001; ONMW26460; T26G1700039/153170HL6. NT2JG1800004; ONMW26463; T26G1700040/155213HL2. NT2JG1800005; ONMW26465; T26G1700041/156108HL3. NT2JG1800006; ONMW26469; T26G1700042/156506HL8. NT2JG1800007; ONMX26485; T26G1700043/157426HL8. NT2JG1800008; ONMX26486; T26G1700044/157506HL7. NT2JG1800009; ONMX26487; T26G1800001/159471HL2. NT2JG1800010; ONMX26504; T26G1800002/160222HL6. NT2QG1900003BH\NT2UG2100001BH; ONMX26505; T26G1800003/160225HL9. NT2UG2100002BH; ONMX26506; T27G15002/ 122818HL8. ONMV26355; ONMX26508; T27G15003/ 124423HL5. ONMV26356; ONMX26509; T27G15004/126534HL7. ONMV26357; ONMX26522; T27G15005/126820HL0. ONMV26358; ONMX26523; T27G15006/GLX-TAN-03716. ONMV26366; ONMX26528; T27G15008/ 128542HL8. ONMV26373; ONMX26532; T27G15011/130994HL7. ONMV26374; ONMX26533; T27G15012/131277HL6. ONMV26380; ONMX26535; T27G15013/ 131436HL8. ONMW26348; ONMX26540; T27G15014. ONMW26349; ONMX26541; T27G15015. ONMW26379; ONMX26543; T27G1600025/0. ONMW26381; ONMX26545; T27G1600026/140419HL3. ONMW26383; ONMX26546; T27G1600027. ONMW26384; T25G12006-R1; T27G1600028. ONMW26387; T26G15018/130763HL6; T27G1600029. ONMW26390; T26G15019/130992HL1; T27G16016/ 133646HL0. ONMW26391; T26G1600026; T27G16017/ 133897HL9. ONMW26395; T26G1600027/0; T27G16018/ 134286HL4. ONMW26402; T26G1600028/0; T27G16019/ 134674HL1. ONMW26403; T26G1600029/144952HL9; T27G16020/ 134675HL8. Appendix A – Impacted Serial Numbers, STOP USING your system: 600 Series Rotor A Model Serial Numbers. T27G16021/136626HL9; T2AG12010/103BH2; T2BG1600020/ 145513HL8. T27G16022/137922HL1; T2AG12011/ 89154HL9; T2BG1600021/208BH9. T27G16023/ Not available; T2AG12012/ 89666HL2; T2BG1600022/ 146064HL1. T27G1700030/147506HL0; T2AG12013/ 90088HL6; T2BG1600023/ 146187HL0. T27G1700031/ 147750HL4; T2AG12014/103BH2; T2BG1600024/ 146363HL7. T27G1700032/ 148489HL8; T2AG12015/118BH0; T2BG1600025/ 146380HL1. T27G1700033/150126HL1; T2AG13016/ 93885HL2; T2BG1600026/146941HL0. T27G1700034/150126HL1; T2AG13017/ 93885HL2; T2BG1600027/147230HL7. T27G1700035/ 151334HL0; T2AG13018/ 94292HL0; T2BG1600028. T27G1700037; T2AG13019; T2BG1600029/ 147600HL1. T28G15001; T2AG13020/ 96237HL3; T2BG1600030/ 147951HL8. T28G15002; T2AG13021; T2BG1600031/ 148143HL1. T28G15002N1; T2AG13023/132BH1; T2BG1600032/ 148379HL1. T28G1600008; T2AG13024/134BH7; T2BG1600033/ 148956HL6. T28G1600010; T2AG13025/100091HL8; T2BG1600034/ 149490HL5. T28G1600011; T2AG13026; T2BG16006/135185HL7. T28G16003; T2AG13027/ 104213HL4; T2BG16007/191BH7. T28G16004; T2AG13029/141BH2; T2BG1700036/148957HL4. T28G16005/135524HL7; T2AG13031/106485HL6; T2BG1700037. T28G16006/0; T2AG13032/ 107349HL3; T2BG1700038/150128HL7. T28G16007; T2AG14033/ 108281HL7; T2BG1700039/ 151143HL5. T28G1700013; T2AG14034; T2BG1700040/ 150067HL7. T28G1700014; T2AG14035/148BH7; T2BG1700041/153537HL6. T28G1700015; T2AG14036/112879HL2; T2BG1700042/ 155272HL8. T28G1700016; T2AG14037/112809HL9; T2BG1700043/ 155484HL9. T28G1700017; T2AG14038/0; T2BG1700044/ Not available. T28G1700018; T2AG14039/115651HL2; T2BG1700045/ 155928HL5. T28G1700019; T2AG14040/ 119081HL8; T2BG1700046/ DDRW34029. T28G1700020; T2AG15041; T2BG1700047/ 156900HL3. T28G1700021; T2BG11001/ 79573HL2; T2BG1700048/ 157050HL6. T28G1700022; T2BG12002/319938HM7; T2BG1700049/ 158144HL6. T28G1800002/159885HL3; T2BG12003/92532HL1; T2BG1700050/ 246BH9. T28G1800003/160874HL4; T2BG13004; T2BG1700051/ 158839HL1. T28G1800005/164844HL3; T2BG14005; T2BG1700052/ 158907HL6. T28G1800006/165131HL4; T2BG1600008/194BH1; T2BG1800001/ 159887HL9. T28G1800007/166055HL4; T2BG1600010; T2BG1800002/ Not available. T28G1800008/166378HL0; T2BG1600011; T2BG1800003/ 23214. T28G1800023/159111HL4; T2BG1600012; T2BG1800005/253BH5. T2AG12003/ 79574HL0; T2BG1600013/141857HL3; T2BG1800006/165346HL8. T2AG12004/22044; T2BG1600014/142225HL2; T2BG1800007/ Not available. T2AG12005/93BH5; T2BG1600015; T2BG1800008/261BH8. T2AG12006/86884HL4; T2BG1600016; T2CG11001/84BH4. T2AG12007; T2BG1600017; T2CG11002. T2AG12008/ 87697HL9; T2BG1600018/ 144093HL2; T2CG12003/116BH4. T2AG12009/ 88227HL4; T2BG1600019/ 144873HL7; T2CG12004. Appendix A – Impacted Serial Numbers, STOP USING your system: 600 Series Rotor A Model Serial Numbers. T2CG13005/125BH5. T2CG13006/127BH1. T2CG14007. T2CG15008/161BH0. T2CG15009/163BH6. T2CG15010/166BH9. T2CG15012/170BH1. T2CG15013/176BH8. T2CG15014/177BH6. T2CG15015/100035454149. T2CG16016/182BH6. T2CG16017/185BH9. T2CG16018/187BH5. T2CG16019/232BH9. T2DG13001/133BH9. T2DG13002. T2DG13003. T2EG15001. T2EG1600001. T2EG1600002/0. T2EG1700003/153301HL7. T2EG1700004. T2EG1700005/153925HL3. T2GG1700002/150684HL9. T2HG1700001. T2HG1700002. T2HG1700003/ 158624HL7. T2HG1800001. T2HG1800002. T2HG1800003. T2HG1800004. T2JG1800001.

Appendix B – Impacted Serial Numbers, continue using your system:

| Infinia and Infinia Hawkeye Model Serial Numbers |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 15006 | 16109 | 16212 | 16309 | 16418 | 16541 | 16652 | 16750 |
| 15007 | 16114 | 16213 | 16310 | 16421 | 16544 | 16654 | 16751 |
| 15013 | 16115 | 16215 | 16312 | 16423 | 16546 | 16656 | 16753 |
| 16001 | 16123 | 16216 | 16313 | 16424 | 16549 | 16657 | 16754 |
| 16004 | 16127 | 16218 | 16315 | 16425 | 16552 | 16658 | 16756 |
| 16005 | 16130 | 16220 | 16318 | 16426 | 16553 | 16659 | 16757 |
| 16006 | 16132 | 16222 | 16323 | 16428 | 16556 | 16660 | 16760 |
| 16007 | 16134 | 16223 | 16324 | 16432 | 16557 | 16662 | 16764 |
| 16009 | 16135 | 16224 | 16329 | 16433 | 16558 | 16663 | 16766 |
| 16015 | 16137 | 16226 | 16331 | 16434 | 16559 | 16664 | 16768 |
| 16019 | 16140 | 16227 | 16333 | 16436 | 16560 | 16665 | 16769 |
| 16021 | 16141 | 16228 | 16335 | 16437 | 16564 | 16668 | 16770 |
| 16022 | 16143 | 16232 | 16336 | 16438 | 16565 | 16670 | 16771 |
| 16026 | 16148 | 16234 | 16338 | 16442 | 16570 | 16673 | 16775 |
| 16027 | 16151 | 16242 | 16341 | 16444 | 16572 | 16676 | 16776 |
| 16028 | 16153 | 16246 | 16347 | 16446 | 16576 | 16677 | 16778 |
| 16030 | 16154 | 16248 | 16348 | 16451 | 16578 | 16678 | 16779 |
| 16031 | 16155 | 16250 | 16350 | 16455 | 16579 | 16680 | 16780 |
| 16033 | 16158 | 16251 | 16352 | 16456 | 16583 | 16683 | 16781 |
| 16037 | 16159 | 16252 | 16353 | 16457 | 16590 | 16692 | 16783 |
| 16038 | 16163 | 16254 | 16355 | 16460 | 16592 | 16694 | 16790 |
| 16039 | 16164 | 16256 | 16357 | 16461 | 16595 | 16695 | 16793 |
| 16043 | 16166 | 16258 | 16360 | 16466 | 16597 | 16699 | 16794 |
| 16044 | 16168 | 16262 | 16361 | 16470 | 16598 | 16700 | 16800 |
| 16045 | 16169 | 16263 | 16362 | 16478 | 16599 | 16702 | 16804 |
| 16048 | 16170 | 16264 | 16370 | 16481 | 16602 | 16703 | 16806 |
| 16049 | 16171 | 16265 | 16371 | 16483 | 16605 | 16705 | 16808 |
| 16050 | 16172 | 16267 | 16373 | 16488 | 16607 | 16706 | 16809 |
| 16051 | 16175 | 16271 | 16374 | 16489 | 16608 | 16707 | 16812 |
| 16052 | 16177 | 16273 | 16376 | 16490 | 16609 | 16709 | 16813 |
| 16054 | 16178 | 16277 | 16382 | 16491 | 16611 | 16712 | 16816 |
| 16057 | 16182 | 16280 | 16383 | 16493 | 16615 | 16714 | 16817 |
| 16059 | 16191 | 16282 | 16386 | 16495 | 16618 | 16715 | 16818 |
| 16060 | 16194 | 16285 | 16390 | 16496 | 16620 | 16716 | 16819 |
| 16063 | 16198 | 16287 | 16399 | 16497 | 16623 | 16717 | 16821 |
| 16064 | 16199 | 16288 | 16401 | 16499 | 16626 | 16720 | 16823 |
| 16070 | 16200 | 16292 | 16402 | 16500 | 16628 | 16725 | 16824 |
| 16073 | 16201 | 16294 | 16404 | 16501 | 16633 | 16729 | 16828 |
| 16082 | 16202 | 16296 | 16408 | 16506 | 16635 | 16734 | 16830 |
| 16083 | 16204 | 16298 | 16409 | 16507 | 16636 | 16736 | 16831 |
| 16084 | 16205 | 16299 | 16412 | 16509 | 16637 | 16738 | 16835 |
| 16091 | 16206 | 16302 | 16413 | 16514 | 16639 | 16739 | 16846 |
| 16094 | 16208 | 16303 | 16414 | 16521 | 16644 | 16747 | 16848 |
| 16099 | 16209 | 16304 | 16415 | 16526 | 16645 | 16748 | 16849 |
| 16107 | 16211 | 16305 | 16417 | 16528 | 16648 | 16749 | 16853 |

Appendix B – Impacted Serial Numbers, continue using your system:

| Infinia and Infinia Hawkeye Model Serial Numbers |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 16856 | 16956 | 17052 | 17141 | 17237 | 17335 | 17429 | 17512 |
| 16857 | 16958 | 17053 | 17144 | 17240 | 17336 | 17431 | 17514 |
| 16858 | 16961 | 17054 | 17146 | 17241 | 17342 | 17433 | 17516 |
| 16859 | 16962 | 17055 | 17149 | 17242 | 17346 | 17436 | 17517 |
| 16860 | 16963 | 17057 | 17150 | 17244 | 17349 | 17438 | 17519 |
| 16861 | 16964 | 17058 | 17153 | 17246 | 17356 | 17439 | 17520 |
| 16863 | 16967 | 17059 | 17155 | 17248 | 17357 | 17441 | 17524 |
| 16864 | 16969 | 17061 | 17156 | 17249 | 17358 | 17442 | 17527 |
| 16865 | 16970 | 17062 | 17160 | 17252 | 17360 | 17443 | 17530 |
| 16866 | 16971 | 17067 | 17162 | 17253 | 17362 | 17444 | 17531 |
| 16869 | 16972 | 17068 | 17163 | 17256 | 17364 | 17446 | 17532 |
| 16870 | 16975 | 17069 | 17165 | 17258 | 17365 | 17447 | 17541 |
| 16871 | 16976 | 17075 | 17166 | 17261 | 17366 | 17448 | 17546 |
| 16877 | 16977 | 17076 | 17167 | 17263 | 17367 | 17450 | 17548 |
| 16882 | 16979 | 17078 | 17169 | 17264 | 17368 | 17451 | 17552 |
| 16883 | 16980 | 17079 | 17174 | 17269 | 17369 | 17452 | 17553 |
| 16886 | 16981 | 17080 | 17175 | 17271 | 17371 | 17454 | 17554 |
| 16887 | 16984 | 17082 | 17176 | 17274 | 17373 | 17456 | 17555 |
| 16888 | 16987 | 17083 | 17177 | 17283 | 17378 | 17457 | 17558 |
| 16891 | 16988 | 17085 | 17179 | 17286 | 17379 | 17465 | 17560 |
| 16892 | 16991 | 17086 | 17182 | 17287 | 17380 | 17466 | 17562 |
| 16897 | 16993 | 17090 | 17184 | 17291 | 17381 | 17468 | 17563 |
| 16900 | 17001 | 17093 | 17187 | 17292 | 17384 | 17470 | 17564 |
| 16902 | 17002 | 17095 | 17188 | 17294 | 17386 | 17471 | 17565 |
| 16910 | 17005 | 17096 | 17189 | 17297 | 17388 | 17472 | 17566 |
| 16912 | 17006 | 17098 | 17190 | 17300 | 17389 | 17478 | 17567 |
| 16913 | 17012 | 17100 | 17191 | 17302 | 17391 | 17481 | 17570 |
| 16916 | 17014 | 17101 | 17195 | 17303 | 17392 | 17483 | 17571 |
| 16920 | 17015 | 17102 | 17196 | 17305 | 17394 | 17484 | 17572 |
| 16921 | 17017 | 17103 | 17201 | 17310 | 17397 | 17485 | 17573 |
| 16923 | 17019 | 17104 | 17203 | 17311 | 17398 | 17486 | 17575 |
| 16925 | 17022 | 17105 | 17204 | 17312 | 17400 | 17487 | 17577 |
| 16927 | 17023 | 17110 | 17206 | 17314 | 17401 | 17488 | 17579 |
| 16929 | 17026 | 17111 | 17207 | 17315 | 17403 | 17489 | 17580 |
| 16931 | 17027 | 17115 | 17211 | 17318 | 17405 | 17491 | 17581 |
| 16935 | 17028 | 17117 | 17218 | 17320 | 17406 | 17497 | 17584 |
| 16940 | 17031 | 17118 | 17220 | 17321 | 17410 | 17498 | 17585 |
| 16942 | 17033 | 17119 | 17222 | 17322 | 17412 | 17500 | 17592 |
| 16943 | 17034 | 17123 | 17223 | 17323 | 17414 | 17501 | 17593 |
| 16946 | 17037 | 17124 | 17224 | 17325 | 17416 | 17502 | 17594 |
| 16948 | 17038 | 17131 | 17229 | 17326 | 17417 | 17504 | 17598 |
| 16949 | 17040 | 17133 | 17230 | 17330 | 17418 | 17505 | 17599 |
| 16950 | 17045 | 17134 | 17231 | 17331 | 17420 | 17506 | 17603 |
| 16953 | 17048 | 17135 | 17235 | 17332 | 17424 | 17508 | 17604 |
| 16955 | 17049 | 17137 | 17236 | 17334 | 17426 | 17510 | 17608 |

Appendix B – Impacted Serial Numbers, continue using your system:

| Infinia and Infinia Hawkeye Model Serial Numbers |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 17613 | 17711 | 17806 | 17909 | 18000 | 18098 | 18181 | 18262 |
| 17615 | 17714 | 17809 | 17911 | 18006 | 18101 | 18183 | 18263 |
| 17618 | 17715 | 17811 | 17913 | 18014 | 18103 | 18184 | 18264 |
| 17619 | 17718 | 17812 | 17914 | 18016 | 18105 | 18185 | 18267 |
| 17622 | 17721 | 17815 | 17917 | 18019 | 18106 | 18188 | 18268 |
| 17627 | 17739 | 17823 | 17924 | 18024 | 18107 | 18190 | 18269 |
| 17628 | 17740 | 17824 | 17925 | 18025 | 18110 | 18193 | 18272 |
| 17632 | 17743 | 17826 | 17929 | 18026 | 18111 | 18195 | 18273 |
| 17634 | 17744 | 17829 | 17934 | 18028 | 18112 | 18197 | 18281 |
| 17636 | 17746 | 17831 | 17936 | 18029 | 18113 | 18198 | 18282 |
| 17637 | 17747 | 17833 | 17938 | 18030 | 18116 | 18201 | 18284 |
| 17638 | 17748 | 17834 | 17939 | 18031 | 18117 | 18204 | 18288 |
| 17642 | 17753 | 17836 | 17940 | 18033 | 18120 | 18205 | 18290 |
| 17643 | 17754 | 17838 | 17941 | 18034 | 18123 | 18207 | 18291 |
| 17650 | 17756 | 17839 | 17944 | 18035 | 18127 | 18208 | 18293 |
| 17653 | 17758 | 17842 | 17947 | 18036 | 18128 | 18210 | 18294 |
| 17654 | 17759 | 17845 | 17948 | 18037 | 18138 | 18211 | 18296 |
| 17655 | 17761 | 17846 | 17950 | 18039 | 18139 | 18212 | 18297 |
| 17656 | 17762 | 17851 | 17951 | 18044 | 18141 | 18216 | 18298 |
| 17660 | 17763 | 17858 | 17952 | 18046 | 18143 | 18217 | 18300 |
| 17663 | 17766 | 17862 | 17953 | 18048 | 18144 | 18218 | 18303 |
| 17664 | 17769 | 17867 | 17954 | 18051 | 18145 | 18219 | 18304 |
| 17667 | 17770 | 17868 | 17956 | 18052 | 18146 | 18220 | 18305 |
| 17668 | 17771 | 17870 | 17957 | 18054 | 18147 | 18221 | 18307 |
| 17674 | 17772 | 17872 | 17960 | 18055 | 18150 | 18222 | 18308 |
| 17675 | 17774 | 17876 | 17966 | 18057 | 18152 | 18224 | 18309 |
| 17676 | 17776 | 17877 | 17968 | 18064 | 18153 | 18225 | 18312 |
| 17677 | 17777 | 17878 | 17971 | 18065 | 18154 | 18228 | 18313 |
| 17679 | 17778 | 17879 | 17973 | 18066 | 18155 | 18229 | 18314 |
| 17680 | 17779 | 17880 | 17974 | 18068 | 18157 | 18230 | 18315 |
| 17683 | 17780 | 17883 | 17975 | 18069 | 18158 | 18231 | 18316 |
| 17684 | 17781 | 17886 | 17978 | 18070 | 18159 | 18232 | 18317 |
| 17687 | 17782 | 17887 | 17981 | 18071 | 18160 | 18235 | 18319 |
| 17688 | 17784 | 17889 | 17982 | 18074 | 18163 | 18236 | 18321 |
| 17692 | 17785 | 17891 | 17983 | 18075 | 18164 | 18238 | 18323 |
| 17693 | 17786 | 17893 | 17984 | 18078 | 18165 | 18244 | 18324 |
| 17696 | 17787 | 17894 | 17985 | 18082 | 18167 | 18245 | 18325 |
| 17698 | 17788 | 17895 | 17987 | 18083 | 18169 | 18246 | 18326 |
| 17699 | 17797 | 17896 | 17991 | 18085 | 18171 | 18250 | 18327 |
| 17702 | 17798 | 17899 | 17992 | 18088 | 18172 | 18251 | 18328 |
| 17703 | 17799 | 17900 | 17994 | 18090 | 18174 | 18254 | 18329 |
| 17704 | 17800 | 17902 | 17995 | 18091 | 18176 | 18256 | 18330 |
| 17705 | 17802 | 17903 | 17997 | 18092 | 18177 | 18258 | 18331 |
| 17707 | 17803 | 17905 | 17998 | 18094 | 18178 | 18260 | 18334 |
| 17708 | 17804 | 17906 | 17999 | 18096 | 18180 | 18261 | 18335 |

Appendix B – Impacted Serial Numbers, continue using your system:

| Infinia and Infinia Hawkeye Model Serial Numbers |  |  |  |  |
| --- | --- | --- | --- | --- |
| 18336 | 18379 | 18429 | 18481 | 18534 |
| 18337 | 18380 | 18432 | 18482 | 18535 |
| 18338 | 18381 | 18434 | 18483 | 18536 |
| 18340 | 18383 | 18437 | 18485 | 18539 |
| 18341 | 18384 | 18438 | 18486 | 18540 |
| 18342 | 18385 | 18443 | 18487 | 18541 |
| 18343 | 18386 | 18444 | 18489 |  |
| 18345 | 18391 | 18445 | 18496 |  |
| 18346 | 18392 | 18446 | 18498 |  |
| 18347 | 18394 | 18449 | 18499 |  |
| 18348 | 18395 | 18450 | 18500 |  |
| 18349 | 18396 | 18451 | 18501 |  |
| 18350 | 18398 | 18452 | 18503 |  |
| 18351 | 18399 | 18453 | 18505 |  |
| 18353 | 18400 | 18454 | 18506 |  |
| 18354 | 18401 | 18456 | 18507 |  |
| 18356 | 18403 | 18459 | 18508 |  |
| 18357 | 18405 | 18460 | 18509 |  |
| 18358 | 18406 | 18461 | 18511 |  |
| 18359 | 18407 | 18462 | 18513 |  |
| 18360 | 18408 | 18464 | 18514 |  |
| 18361 | 18409 | 18465 | 18515 |  |
| 18365 | 18411 | 18466 | 18516 |  |
| 18368 | 18412 | 18467 | 18518 |  |
| 18369 | 18417 | 18470 | 18520 |  |
| 18370 | 18418 | 18471 | 18521 |  |
| 18371 | 18419 | 18472 | 18526 |  |
| 18372 | 18421 | 18473 | 18527 |  |
| 18374 | 18423 | 18474 | 18528 |  |
| 18376 | 18424 | 18477 | 18531 |  |
| 18378 | 18427 | 18479 | 18532 |  |

| Varicam and VG Model Serial Numbers |  |  |  |
| --- | --- | --- | --- |
| 10539 | 10543 | 10553 | 10561 |
| 10571 | 10579 | 10599 | 10602 |
| 10604 | 1003 |

Appendix B – Impacted Serial Numbers, continue using your system:

| 600 Series Rotor A Model Serial Numbers |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 21112 | 22193 | 22363 | 22501 | 22594 | 26109 | 26244 | 26344 |
| 21208 | 22195 | 22365 | 22506 | 22595 | 26113 | 26245 | 26345 |
| 21210 | 22196 | 22372 | 22509 | 22598 | 26114 | 26246 | 26347 |
| 21227 | 22200 | 22373 | 22513 | 22600 | 26116 | 26247 | 26350 |
| 21249 | 22203 | 22374 | 22515 | 22608 | 26118 | 26251 | 26352 |
| 21254 | 22204 | 22381 | 22516 | 22609 | 26120 | 26253 | 26353 |
| 21257 | 22213 | 22382 | 22517 | 24022 | 26122 | 26254 | 26354 |
| 21261 | 22215 | 22384 | 22518 | 24023 | 26125 | 26261 | 27001 |
| 21278 | 22217 | 22390 | 22520 | 24024 | 26126 | 26262 | 27003 |
| 21282 | 22226 | 22395 | 22522 | 24026 | 26131 | 26264 | 27005 |
| 21293 | 22227 | 22396 | 22523 | 24056 | 26133 | 26265 | 27007 |
| 21297 | 22235 | 22397 | 22524 | 24060 | 26139 | 26266 | 27008 |
| 21303 | 22250 | 22399 | 22526 | 24064 | 26140 | 26268 | 27011 |
| 21306 | 22251 | 22400 | 22527 | 24066 | 26146 | 26270 | 27016 |
| 21308 | 22253 | 22409 | 22528 | 24070 | 26147 | 26273 | 27018 |
| 21321 | 22255 | 22410 | 22529 | 25204 | 26148 | 26279 | 27025 |
| 21341 | 22258 | 22414 | 22530 | 26013 | 26149 | 26280 | 27029 |
| 21343 | 22263 | 22417 | 22532 | 26015 | 26157 | 26281 | 27031 |
| 21350 | 22264 | 22422 | 22534 | 26025 | 26159 | 26282 | 27041 |
| 21372 | 22269 | 22424 | 22535 | 26030 | 26161 | 26283 | 27048 |
| 21373 | 22270 | 22425 | 22536 | 26033 | 26162 | 26284 | 27049 |
| 21374 | 22271 | 22428 | 22541 | 26034 | 26165 | 26285 | 27051 |
| 21394 | 22272 | 22429 | 22542 | 26035 | 26166 | 26286 | 27058 |
| 22104 | 22275 | 22432 | 22545 | 26036 | 26168 | 26291 | 27060 |
| 22107 | 22282 | 22434 | 22546 | 26040 | 26174 | 26293 | 27064 |
| 22108 | 22286 | 22435 | 22550 | 26041 | 26179 | 26294 | 27065 |
| 22109 | 22292 | 22440 | 22552 | 26043 | 26181 | 26296 | 27071 |
| 22113 | 22297 | 22441 | 22558 | 26044 | 26185 | 26299 | 27075 |
| 22126 | 22300 | 22445 | 22559 | 26049 | 26190 | 26302 | 27077 |
| 22137 | 22307 | 22446 | 22560 | 26050 | 26195 | 26304 | 27080 |
| 22139 | 22312 | 22450 | 22564 | 26056 | 26196 | 26309 | 27089 |
| 22140 | 22318 | 22455 | 22565 | 26057 | 26197 | 26310 | 27092 |
| 22143 | 22321 | 22459 | 22567 | 26065 | 26204 | 26312 | 27098 |
| 22145 | 22324 | 22460 | 22569 | 26080 | 26206 | 26315 | 27100 |
| 22147 | 22328 | 22461 | 22570 | 26083 | 26209 | 26317 | 27102 |
| 22148 | 22331 | 22462 | 22572 | 26088 | 26216 | 26320 | 27103 |
| 22154 | 22334 | 22467 | 22573 | 26089 | 26218 | 26321 | 27110 |
| 22159 | 22337 | 22472 | 22579 | 26090 | 26220 | 26322 | 27115 |
| 22163 | 22339 | 22485 | 22581 | 26091 | 26221 | 26330 | 27117 |
| 22174 | 22341 | 22486 | 22584 | 26093 | 26223 | 26333 | 27118 |
| 22178 | 22346 | 22487 | 22586 | 26098 | 26227 | 26334 | 27119 |
| 22186 | 22350 | 22492 | 22588 | 26099 | 26229 | 26337 | 27121 |
| 22188 | 22351 | 22495 | 22590 | 26101 | 26231 | 26339 | 27124 |
| 22190 | 22360 | 22498 | 22591 | 26102 | 26235 | 26341 | 27126 |
| 22192 | 22361 | 22500 | 22592 | 26103 | 26243 | 26343 | 27127 |

Appendix B – Impacted Serial Numbers, continue using your system:

| 600 Series Rotor A Model Serial Numbers |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| 27129 | 630V22670 | 630W22740 | 630W22798 | 630W22863 | DDRW34010 |
| 27131 | 630V22671 | 630W22743 | 630W22800 | 630W22864 | DDRW34013 |
| 27133 | 630V22672 | 630W22744 | 630W22801 | 630W22865 | DDRW34015 |
| 28001 | 630V22674 | 630W22745 | 630W22804 | 630X22870 | DDRW34016 |
| 28003 | 630V22675 | 630W22748 | 630W22806 | 630X22871 | DDRW34018 |
| 28004 | 630V22677 | 630W22749 | 630W22807 | 630X22872 | DDRW34020 |
| 630V22612 | 630V22686 | 630W22750 | 630W22808 | 630X22873 | DDRW34023 |
| 630V22615 | 630V22688 | 630W22751 | 630W22809 | 630X22874 | DDRW34024 |
| 630V22616 | 630V22689 | 630W22752 | 630W22811 | 630X22875 | DDRW34025 |
| 630V22617 | 630V22690 | 630W22754 | 630W22812 | 630X22876 | DDRW34028 |
| 630V22621 | 630V22691 | 630W22755 | 630W22814 | 630X22878 | DDRW34031 |
| 630V22622 | 630V22694 | 630W22756 | 630W22815 | 630X22879 | DDRW34032 |
| 630V22625 | 630V22695 | 630W22757 | 630W22817 | 630X22880 | DDRW34033 |
| 630V22626 | 630V22696 | 630W22758 | 630W22818 | 630X22883 | DDRX34038 |
| 630V22627 | 630V22699 | 630W22759 | 630W22819 | 630X22889 | DDRX34039 |
| 630V22628 | 630V22700 | 630W22760 | 630W22820 | 630X22890 | DDRX34040 |
| 630V22629 | 630V22701 | 630W22761 | 630W22821 | 630X22891 | DDRX34041 |
| 630V22630 | 630V22702 | 630W22762 | 630W22822 | 630X22892 | DDRX34042 |
| 630V22631 | 630V22703 | 630W22763 | 630W22823 | 630X22894 | DDRX34044 |
| 630V22632 | 630V22706 | 630W22765 | 630W22824 | 630X22896 | DDRX34045 |
| 630V22633 | 630V22707 | 630W22767 | 630W22825 | 630X22898 | DDRX34046 |
| 630V22637 | 630V22708 | 630W22768 | 630W22827 | 630X22899 | DDRX34047 |
| 630V22638 | 630V22709 | 630W22769 | 630W22830 | 630X22902 | DDRX34049 |
| 630V22639 | 630V22710 | 630W22772 | 630W22831 | 630X22905 | DDRX34053 |
| 630V22640 | 630V22712 | 630W22773 | 630W22832 | 630X22906 | DDRX34055 |
| 630V22641 | 630V22714 | 630W22774 | 630W22833 | 630X22907 | DDRX34056 |
| 630V22642 | 630V22715 | 630W22775 | 630W22835 | 630X22910 | DDRX34058 |
| 630V22644 | 630V22716 | 630W22776 | 630W22837 | 630X22914 | DESV28013 |
| 630V22647 | 630V22717 | 630W22777 | 630W22838 | 630X22916 | DESV28019 |
| 630V22648 | 630V22719 | 630W22778 | 630W22839 | 630X22917 | DESV28020 |
| 630V22651 | 630V22722 | 630W22779 | 630W22840 | 630X22919 | DESV28021 |
| 630V22652 | 630W22724 | 630W22782 | 630W22842 | 630X22921 | DESV28022 |
| 630V22653 | 630W22725 | 630W22783 | 630W22843 | 630X22922 | DESW28024 |
| 630V22654 | 630W22726 | 630W22785 | 630W22845 | 630X22923 | DESW28025 |
| 630V22655 | 630W22727 | 630W22786 | 630W22846 | 630X22924 | DESW28029 |
| 630V22656 | 630W22728 | 630W22787 | 630W22848 | 630X22925 | DRGX56002 |
| 630V22658 | 630W22729 | 630W22788 | 630W22849 | 670V21426 | DRGX56012 |
| 630V22659 | 630W22730 | 630W22789 | 630W22850 | CZTV29005 | ONMV26359 |
| 630V22660 | 630W22732 | 630W22790 | 630W22851 | CZTW29007 | ONMV26360 |
| 630V22662 | 630W22733 | 630W22792 | 630W22856 | CZTW29009 | ONMV26361 |
| 630V22663 | 630W22734 | 630W22793 | 630W22857 | DDRW34003 | ONMV26362 |
| 630V22664 | 630W22735 | 630W22794 | 630W22858 | DDRW34004 | ONMV26363 |
| 630V22665 | 630W22736 | 630W22795 | 630W22859 | DDRW34005 | ONMV26364 |
| 630V22666 | 630W22738 | 630W22796 | 630W22860 | DDRW34006 | ONMV26365 |
| 630V22668 | 630W22739 | 630W22797 | 630W22862 | DDRW34007 | ONMV26367 |

Appendix B – Impacted Serial Numbers, continue using your system:

| 600 Series Rotor A Model Serial Numbers |  |  |
| --- | --- | --- |
| ONMV26368 | ONMW26435 | ONMX26499 |
| ONMV26369 | ONMW26436 | ONMX26500 |
| ONMV26370 | ONMW26437 | ONMX26502 |
| ONMV26371 | ONMW26439 | ONMX26503 |
| ONMV26372 | ONMW26441 | ONMX26507 |
| ONMV26375 | ONMW26442 | ONMX26510 |
| ONMV26376 | ONMW26443 | ONMX26511 |
| ONMV26378 | ONMW26444 | ONMX26512 |
| ONMW26382 | ONMW26446 | ONMX26513 |
| ONMW26385 | ONMW26447 | ONMX26514 |
| ONMW26386 | ONMW26451 | ONMX26515 |
| ONMW26388 | ONMW26452 | ONMX26516 |
| ONMW26389 | ONMW26454 | ONMX26517 |
| ONMW26392 | ONMW26457 | ONMX26518 |
| ONMW26393 | ONMW26458 | ONMX26519 |
| ONMW26394 | ONMW26459 | ONMX26520 |
| ONMW26396 | ONMW26461 | ONMX26521 |
| ONMW26397 | ONMW26462 | ONMX26524 |
| ONMW26398 | ONMW26464 | ONMX26525 |
| ONMW26399 | ONMW26466 | ONMX26526 |
| ONMW26400 | ONMW26468 | ONMX26527 |
| ONMW26401 | ONMW26470 | ONMX26529 |
| ONMW26404 | ONMW26471 | ONMX26530 |
| ONMW26405 | ONMW26472 | ONMX26531 |
| ONMW26407 | ONMW26473 | ONMX26534 |
| ONMW26408 | ONMW26474 | ONMX26536 |
| ONMW26409 | ONMW26501 | ONMX26538 |
| ONMW26410 | ONMX26476 | ONMX26539 |
| ONMW26411 | ONMX26477 | ONMX26542 |
| ONMW26413 | ONMX26479 | ONMX26544 |
| ONMW26414 | ONMX26480 |  |
| ONMW26415 | ONMX26481 |  |
| ONMW26417 | ONMX26482 |  |
| ONMW26418 | ONMX26483 |  |
| ONMW26420 | ONMX26484 |  |
| ONMW26422 | ONMX26488 |  |
| ONMW26423 | ONMX26489 |  |
| ONMW26424 | ONMX26490 |  |
| ONMW26425 | ONMX26491 |  |
| ONMW26426 | ONMX26492 |  |
| ONMW26428 | ONMX26493 |  |
| ONMW26429 | ONMX26494 |  |
| ONMW26430 | ONMX26495 |  |
| ONMW26431 | ONMX26496 |  |
| ONMW26433 | ONMX26497 |  |